TMS - Training
General Overview of the TMS system
TMS hardware
TMS Software
Cycle Parameters (State/Phase tables)
FPGA Firmware packaging
Testing
Troubleshooting

---
TMS - Introduction
Particle trajectory measurement system for the CERN Proton Synchrotron
40 Pick-up Channels - 120 analogue channels sampled at 125MHz, 14 bits
FPGA based system processes this data in real-time
Captures and processes about 15 billion samples per second
Network Data access at approx 65 MBytes per second
256 MBytes of memory per pick-up channel
Gigabit Ethernet interface used internally and externally
Modular and scalable system
Software open source Linux based
FPGA firmware written in VHDL

---
TMS Overview

---
External signals

---
TMS – Hardware

---
TMS – Hardware Installed in Rack
TMS System in rack at CERN
System Controller is black unit above racks.
Network switch is installed behind rack.
Only a few timing and analogue connection have been made.

---
TMS - Hardware
System Controller – Dual Intel Xeon PC architecture system with 2 GBytes of RAM, dual Gigabit Ethernet ports and dual SATA disks in a RAID1 configuration. Boots from hard disk system. Has spare PCI slots for post processing modules or extra network ports.
Module Controller – Intel Duo Core based cPCI controller with 1GByte of RAM and 3 Gigabit Ethernet ports. Boots over Ethernet from System controller. Has PMC slots for extra post processing modules.
PUPE Boards – Xilinx Virtex-4 FX100 based, 1GByte of SDRAM, 2 Gigabit Ethernet ports, 9 x 14bit ADC's, 13 Digital I/O lines, 1 Digital clock input.
Rack systems – Power supply for cPCI systems. 4 separate 8 slot card frames, 1 card frame as a spare.

---
Hardware Installation and Setup
System Controller – No special set-up, BIOS at defaults. One Ethernet port to LAN and the other to the TMS private network switch.
Module Controller – Some BIOS changes, can be set-up using VGA monitor and keyboard plugged into the special breakout lead. Install in far right slots in cPCI rack with transition module at the rear. Network interface Eth2 to TMS switch.
PUPE Boards – Installed in the slots leaving a spare slot next to the power supply and a free slot next to the Module Controller to reduce ADC noise. The Master PUPE should in the far right slot. This has the extra digital timing panel installed. The timing bus cable should connect all of the PUPE's together and the PUPE farthest away from the Master PUPE should have the timing bus termination jumper installed.
TmsMaintenance manual gives full details.

---
PUPE Board

---
Hardware Maintenance
Change hard disks – one every three years. The disk system runs as a RAID 1 system, so one disk can be swapped without having to re-install the software.
Change real-time clock batteries – once every 5 years. The System Controller and Module controllers have these.
Clean the airflow paths – once every 3 years.
Replace the cooling fans in the rack once every 5 years.
Replace the cooling fans in the System Controller every 5 years.
The System controllers and PUPE boards have on-board thermal monitoring. This can be used to check that the fans are working correctly.

---
Hardware Trouble shooting
System Controller – Spare controller can be used. The IPMI interface can be used to diagnose boot problems or a VGA Monitor and keyboard can be attached. Normal PC architecture
Module Controllers – Spare controller can be used. VGA Monitor and keyboard can be attached via breakout lead to diagnose boot problems.
PUPE Boards – Spare PUPE boards can be used. The TMS system performs basic PUPE board tests and gives status information on the power supply voltages and temperature
PUPE boards are identified by module and slot position. A spare PUPE in the spare rack can be used to replace a faulty PUPE without a power cycle.
Power Supply's – These will show a red light if there is a problem. There is a spare power supply in the system.

---
TMS - Software
Operating System is Linux, based on Fedora Core 6 distribution.
Module controller has a very small, network boot Linux system based on the Busybox utility.
TMS software is predominantly written in 'C++' in an object orientated style.
TMS software is multi-threaded.
The GNU software development tool-set is used for development.
The TMS Server has the complete development environment installed.
Two main programs: TmsServer and TmsPuServer.

---
TMS Software Structure

---
File structure

---
Software Configuration
The System Controllers Linux configuration is standard. Limited package installation. tmsSetup sets-up default Network config.
The supplied install DVD has a Kickstart file to install and configure a system from bare metal. Only the network information and Cycle Parameter information needs to be configured.
Users can be added to the system if required.
The TMS software is configured with 5 main files.
/etc/tmsServer.conf – This configures the master TmsServer program.
/usr/tms/rootfs-[1234]/etc/tmsPuServer.conf – These configure the individual TmsPuServer programs running on the module controllers.
The Cycle parameters data is in /usr/tms/stateTables

---
TmsServer.conf

---
TmsPuServer.conf

---
Module Controller Setup
The Module controller boots from the System Controller.
The kernel is loaded from /tftpboot/tms-mcsys using TFTP
The root file system is in: /usr/tms/rootfs-[1234]
There is a utility named “tmsSetupModuleController”. This will copy the template /usr/tms/rootfs into the appropriate root file system directory and optionally add the Ethernet MAC address to the System Controllers /etc/dhcpd.conf file.
It takes the module controller number as the first argument and optionally the Ethernet MAC address as the second argument.
Note that a default tmsPuServer.conf file will be installed from /usr/tms/rootfs. This will probably need editing, especially for Module Controller 3 and 4.

---
Module Controllers System
Fetches network information using DHCP from system controller.
Fetches the Linux kernel and initial RAM disk root file system using TFTP from system controller's /tftpboot/tms-mcsys.
Mounts the /usr/tms/rootfs[1234] file system read only using NFS from system controller.
Mounts local RAM file systems on /tmp and other appropriate places. Startup file /etc/init.d/rcS
Mounts the /usr/tms directory read only using NFS from system controller.
Mounts the /data directory read/write using NFS from system controller.
Busybox is used to implement most of the system programs and files as well as some shared libraries from a Fedora Core 6 system.

---
Software maintenance
The Linux system needs no normal system maintenance.
The TMS configuration files and Cycle Parameter files need to be backed up. The TMS script program “tmsBackup” will create a simple tar archive of the TMS configuration and Cycle Parameter files.
Any additional files added to the system, such as user's home directories need to be backed up.
All software is packaged as RPM packages, including the TMS software. This allows the “yum” and “rpm” tools to manage updating individual software packages.
System log messages in /var/log/messages.

---
TMS Software Interfacing
The TMS system implements a TCP/IP socket based RPC interface for control and data access.
The RPC system is based on the BEAM BOAP object access system. This implements an efficient binary communications system for performance.
A 'C++' API library, libTms, is provided that can be ported to different systems.
Two separate object interfaces: TmsControl and TmsService.
Multiple clients can access the system simultaneously.
Also implements an asynchronous event system.

---
Data Client applications using the TMS API
The TmsService object provides a few simple RPC functions.
Provides information functions for a given cycle type or an individual cycle number.
Returns data from the system.
Supports raw pick-up data (integrated per bunch).
Supports averaged pick-up data (integrated per ms)
Performance is around 65 MBytes/second across a Gigabit network interface.
Data comes from the PUPE memory. This is sufficient for 2 to 3 PS cycles worth of data.
Data bandwidth restricts the amount of data that can be returned to the user.
The main function call is: getData(DataInfo info, Data& data)

---
TMS GetData call

---
TMS GetCycleInformation call

---
TMS GetCycleTypeInformation call

---
Control client applications using the TMS API
SetNextCycle call needed to provide the TMS system with information on the next PS cycle. Needs to be called at least 30ms before the START_CYCLE event. Best time at CYCLE_STOP
Cycle Parameter table management
Diagnostics
System Testing
System Status
System Statistics

---
TMS Software Time Line

---
TMS Software Structure

---
TMS Errors

---
TMS Web Access
The TMS Server presents a simple web access system
Able to view system status
Able to view statistics
Able to get data from the system
Able to display simple graphs.
Could be easily extended.

---
TMS User Programs
tmsControl: Command line control application.
tmsControlGui: GUI Command line application
tmsTestData: Test application
  - “tmsTestData -simdata”
  - “tmsTestData -test all -check -cont”
tmsStateGen: Cycle parameter generator
tmsSigGen: Test signal generator
tmsRestart: Restarts tmsServer and tmsPuServer programs

---
TMS internal software development
All of the TMS software is available in source code form.
The TMS software can be built on the TMS system or another Fedora Core 6 Linux system.
Split into the following Modules:
  - Tms – The main system software
  - Tms-sys – System Controller configuration
  - Tms-mcsys – Module Controller system
  - Tms-fpga – FPGA firmware
  - Tms-doc – system documentation
The tms-full-src-<VERSION>.tar.gz archive contains the full source code.

---
TMS Software Libraries
libBeam.a: BEAM Object API. Includes basic String, List and Array classes as well as the BOAP RPC system.
libBDebug.a: BEAM Debug utilities include crash backtrace system.
libTms.a: Main TMS API library. Includes client and server side BOAP interface objects as well as CycleParameter table generation and management classes.
libadmxrc2.so: Alpha Data ADMXRC interface library. Uses the admxrc2 Linux kernel driver to communicate with the PUPE boards.

---
Building and packaging main software
Uses the make system
Overall make configuration in “Makefile.config”. Includes version number.
  - “make clean” - Cleans the software tree
  - “make” - Builds the software tree
  - “make rpm” - Builds the RPM package
  - “make rpmInstall” - Installs the RPM's in the packages directory
You can also use make install, as root, to install the software directly without packaging it first.
Note tms-mcsys uses the kernel and drivers (including Admxrc2) from the build system.

---
TMS Cycle Parameter Tables
Contains FPGA state and phase table parameters describing a complete PS machine cycle.
Information on what to do on events and internal PLL tables.
Cycle is split into Cycle Periods: Calibration, Event0, Event1, ...
The FPGA hardware can acquire a complete set of data for a cycle with no software intervention.
The TMS Server keeps a library of Cycle Parameter tables in ASCII files indexed by a cycle type string.
The Cycle Parameter tables are passed to all Module controllers which store there contents in internal data structures. The same information is sent to all PUPE channels.
The Cycle Parameters are loaded into the FPGA's on the CYCLE_STOP event or on the setNextCycle() call whichever is later.

---
FPGA State tables
Set of 16 possible states
Change of state on timing event or a delay of 16 FREF periods.
Set of control bits for each state
State 14 is error state – interrupt generated
State 15 is stopped state
Each state has a separate phase table (512 bytes)

---
TMS State Tables – Control bits

---
State/Phase Table Switching

---
FPGA State Tables - Events

---
FPGA Phase tables

---
How to generate Cycle Parameter's
Cycle Parameters are stored in ASCII files on the TMS Server.
New or updated sets of parameters can be uploaded using the TMS API. The TmsServer program will send these out to all module controllers.
The tmsControl and tmsControlGui programs can read an ASCII Cycle Parameter file and send it to the TMS System using the TMS API. So the user can generate a set of Cycle Parameters in an ASCII file and upload this to the TMS server.
The tmsControlGui program has a simple high level Cycle Parameter editor built in.
The tmsStateGen program creates a set of simple test Cycle Parameter's. It can be extended to support other Cycle types.
The TMS API library has support for reading and writing the Cycle Parameter files and creating them based on a high level definition.

---
TMS FPGA Firmware
The TMS FPGA Firmware is packaged in RPM format
Written in VHDL
Build environment is Xilinx Foundation Express
To package:
  - Build bit file. Make sure internal version number is updated.
  - Copy bit file to tms-fpga directory with appropriate version number
  - Symbolically link this bit file with tms-fpga.bit
  - Run “make rpm” to build package
  - Run “make rpmInstall” to install the package in packages
Make sure bit file does not cause overheating of the FPGA
The cern-tms-rel_<VERSION>.zip archive contains the source code.

---
FPGA VHDL Synthesis
Synthesis is achieved using ISE tools (9.2)
Build uses make file (makefile_tms_pupe or makefile_tms_pupe_inc)
The XST project file and build scripts are: tms_pupe_xst.scr and tms_pupe_xst.prj
The project can be built from the command line using nmake -f makefile_tms_pupe
Synthesis scripts and files are located in the synthesis directory. The UCF file contains placement constraints for the SDRAM controller block memories. Without these constraints memory read errors can occur.
3 main directories: tms-processing contains the signal processing code core, pupe-wrapper contains the interfacing code, ddr2_memory_interface contains the DDR2 memory interfaces.

---
FPGA Block Diagram

---
FPGA PupeAPI
Register and shared memory Interface
4 MByte PCI window
System Control Registers: controlling the memory paging and interrupts
Application Registers: for use by the PUPE modules
2 MB Memory window: for accessing SDRAM or Block RAM banks

---
FPGA Memory Map

---
FPGA Application Registers

---
FPGA Shared	Memory

---
FPGA Interrupts

---
TMS Testing
The system has a number of features for testing:
  - Internal API's test function can to perform a basic system test.
  - Diagnostics API functions to capture important internal FPGA signals.
  - Software simulation of timing signals.
  - Generation of Sigma, DeltaX, DeltaY and FRef input signals from SDRAM based signal generator.
  - Messages in /var/log/messages.
  - Low level program debug arguments.
  - TMS test signal generator.

---
Testing using simulated timing and data
Useful test method when no live data is available.
Timing signals generated in software and applied to Master PUPE in each sub-rack.
Data source comes from PUPE SDRAM instead of ADC input.
Can be setup manually using the tmsControlGui's “Pupe Simulation” tab.
The tmsTestData's “-simdata” option allows the complete system to be set to simulation mode.
Note that the “SimulateNextCycle:” field in the /etc/tmsServer.conf file needs to be set to 1 to simulate the setNextCycle call. External systems should be disabled from calling this call.
The tmsTestData program can be used to run soak tests on the data. “tmsTestData -test all -check -cont localhost”

---
Testing on a live system
The Test API function can use used to perform an overall test.
The Status API function can be used to find the status of the system, including the voltages and temperatures of the PUPE FPGA boards.
The Statistics API function keeps a count of errors that have occurred.
The Diagnostics capture function can be used to look at the internal FPGA signals.
The /var/log/messages log can be viewed for any warning or error messages.
The API's errorEvent can be monitored for errors.

---
TMS Trouble shooting
Hardware trouble shooting has been given earlier
Use the systems Test API function via the tmsControlGui program or the web interface.
If the systems Test API function is not working check that the “tmsServer” program is running on the system controller and that the “tmsPuServer” programs are running on the module controllers.
Check the /var/log/messages file for any errors listed.
The system can be restarted using the “tmsRestart” command.
If the “tmsServer” or any “tmsPuServer” programs crash, a backtrace will be listed in the /var/log/messages file.
The tmsServer or tmsPuServer programs can be started manually using the “-f” and “-d 0x03” flags to run them in the foreground and display debug messages. See the manuals on the programs for the “-d” options available.

---
Multiple System Controller Support
Two TmsServers: 192.168.100.1 and 192.168.100.2
Both systems can be setup identically apart from:
  - DHCP disabled on second server
  - TmsServer disabled on second server
  - tmsServer.conf and tmsPuServer.conf files set to use second server.
Second Server could be configured to manage spare module and 3 PUPE boards.
  - Server would supply DHCP information for spare module controller.
  - tmsServer.conf and tmsPuServer.conf files set to use second server.

---
TMS Future Development
Improving the FPGA algorithms including: BLR, PLL.
Improving Software functionality.
Adding data post processing algorithms.
Adding support for linked PLL's taking Sigma from three channels.
Adding extra software API functions.
Increasing data bandwidth by using PUPE Ethernet interfaces.
Adding FPGA based post processing.
Could allow both System Controllers to function with automatic handover.
Could reduce ADC noise. Internal panels. Move power supplies.

---
TMS – Further Information
Further information is available on the support website at: http://portal.beam.ltd/support/cern